
Smart Medication Adherence Phone Case
Section A05, Smart Medication Adherence Phone Case Team

ECE 4012 Senior Design Project

Project Advisor, Dr. Maysam Ghovanloo 

 and Dr. Kevin Murnane

Kedeng Pan, kedengpan@gatech.edu 
Xiannan Di, dixiannanben@gatech.edu

Xingyuan Zhu, xingyuan.zhu@gatech.edu

Yuhao Lin, ylin354@gatech.edu 
Zhipeng Shao, ryanshao@gatech.edu
Submitted

May 03, 2018
Executive Summary
Adherence to medical product has been proved efficient to user’s health. However, the problem of a considerable number of patients not following their prescriptions regularly persists. Regarding to the issue, a software-hardware integrated solution utilizing a portable and cost-efficient medical dispensing device, the Smart Medication Adherence Phone Case, is being proposed. The device is a small pill tray integrated as a phone case and attached to the back of an Android smartphone. It can automatically dispense pills via the Bluetooth Low Energy control from the application and record the medical taking behavior back to the smartphone application. Once filled with medical adherence information either by patients themselves or their doctors, the smartphone application will remind the user to take medicines at proper time by sound and vibration notifications. The mechanism in the case is designed to be able to dispense pills piece by piece automatically. The magnetic position sensor senses the pill-dropping accuracy and auto-correct misbehaviors caused by fabrication imperfectness. Linked with medical center and family members, the project will be a great assistant for elder people in case of emergency. With considerations of parts and services being used, the cost of this device would be lowered below $30. The final outcome of this project will be a fully functional prototype of the Smart Medication Adherence Phone Case device integrated with an Android application. 
Table of Contents
Executive Summary
2
1 Introduction
4
1.1 Objective
4
1.2 Motivation
5
1.3 Background
6
2 Project Description and Goals
8
3 Technical Specifications & Verification
10
4 Design Approach and Details
14
5 Schedule, Tasks, and Milestones
20
6 Final Project Demonstration
21
7 Marketing and Cost Analysis
23
7.1 Marketing Analysis
23
7.2 Cost Analysis
23
8 Conclusion
28
9 References
29
Appendices 
Smart Medication Adherence Phone Case
1.
Introduction
Smart Medication Adherence Phone Case is a convenient and cost-efficient solution that dispenses medications to patients and monitors their medication adherence conditions. The solution consists of a pill-dispensing phone case and an Android application.  
1.1 Objective
The team prototypes a medication dispensing and monitoring system. Through this system, a patient can retrieve dispensed medications from the pill-dispensing phone case, and refill medications to it upon approval from the Android app. Using the sensors built-in on the Android smartphone, the Android app is biometrically secure and accessible: it requires and recognizes a patient’s fingerprint to allow further access. The app enables patients to input the dose and time for his or her medications, set and receive reminders, and monitor individual medication adherence data. The phone case communicates with the Android app via Bluetooth Low Energy (BLE) protocol.

[image: image1.jpg]Andriod Smariphone
fe——— Caietootn §——»
ispensing Phone Case

Medication Adherene app




Figure 1. System overview block diagram of Smart Medication Adherence Phone Case.
1.2 Motivation
Pharmaceutical compliance is a huge medical issue nowadays. Research shows that “between $100 and $300 billion of avoidable annual healthcare costs” attribute to patients to taking their medications on time [1]. The project aims at building a low-cost and portable medication adherence system that separates from other expensive, clumsy and bulky pill-dispensing products on the market. The team decides to take advantage of the popularity of today’s smartphones and reuse their built-in sensors. This approach not only significantly reduces the cost, but also trims down the physical size of the pill-dispensing phone case. The final product targets at patients, especially seniors, who need to adhere to their medications on a daily basis with specific types of the pill. The case appears as a convenient add-on gadget to their smartphones. As patients carry their smartphone around, they can still easily adhere to their medications regardless of their physical locations.
1.3 Background
Many commercial products that aid patients’ medication adherence are already on the market. However, extensive research conducted by the team shows that redundant logging, expensive pricing, and large physical size are usually the drawbacks for the existing products.

Smart Wireless Pill Bottles

AdhereTech builds a “smart” bottle that senses the amount of medication inside the bottle and updates the value to the cloud server [2]. Furthermore, the OS software API offered by the company leads to an integration of hardware, mobile software and medication platforms. The retail fee is about two dollars each. Another type of pillbox, called uBox, is owned by Abiogenix [3]. It connects to schedules and calendars stored on user mobile devices. At the beginning of dosing time, it unlocks itself and automatically supplies desired medicine. This feature prevents accidental overdosing. The product does not retail; instead, the company cooperates with insurance company, clinical center and pharmacy to supply a complete solution that covers from the beginning of prescription to the end of recovery. Compared to the proposed solution by the team, AdhereTech’s smart pill bottle is as conveniently portable, but there is no mobile application with which a patient can log and track their adherence data.

Philips Medication Dispensing Service
Philips’ Medication Dispensing Service is a subscription-based, pill-dispensing machine. It can hold up to 60 cups of medication and dispense medications according to schedule [4]. The overall size of the Philips dispenser is comparable to that of a water boiler. The large size is due to a spinner mechanism inside the dispenser, and a reservoir bank that holds medications not taken on time, preventing overdose. The Dispenser requires connection with a home telephone in order to send out notification to caregivers when patients miss their medications. There is a one-time installation charge of $100 and a monthly $59.95 subscription fee. Compared to the proposed solution by the team, Philips pill-dispenser is immobile and unaffordable. Despite being a complete product, Philips Dispenser is not a viable solution for the majority of patients, such as seniors, who need strict medical adherence.
2.
Project Description and Goals
The final solution consists of two major components: a Smart Pill-Dispensing Phone Case, and an Android application.
The Pill-Dispensing Phone Case has a dimension similar to an eyeglass case. The exterior portion of the case is laser-cutted with transparent acrylic boards. Inside the container it includes a two-layer design that separates the microcontroller and pill containers. In the upper layer, there are two 3D printed dispensing channels that are also able to hold up tp 20 pills per channel. On the other hand, the bottom layers consist the nRF52 microcontroller board by Adafruit that enables Bluetooth Low Energy communication with smartphone, a rechargeable lithium ion battery power supply and necessary wiring.
The mobile app is developed on the Android platform. The Android app integrates a user-friendly GUI through which patients are notified timely for their medication and their adherence data are recorded.

General Features

· Phone Case that dispenses medications and records medication adherence data

·  Physical components required to prototype the Phone Case is around $80

Pill-Dispensing Phone Case

· Holds and dispenses medications to patients on schedule

· Communicates with Android app by BLE

Android App

· Notifies patients to take medications 
· Monitors hardware status (battery level, number of pills left etc.) 
· Biometrically controlled through fingerprint sensor on smartphone
· Record pill-taking behavior of patients

3. Technical Specifications & Verification
3.1 Adafruit Feather nRF52 Bluefruit LE
Table 1 contains product specification for Bluetooth Low Energy board with a native-Bluetooth chip, the nRF52832, implemented on the phone case.

	Table 1. Adafruit Feather nRF52 Bluefruit LE Specification

Item

Specification

Core

ARM® Cortex®-M4 32-bit processor with FPU, 64 MHz

Power

1.7 V–3.3 V voltage input; up to +4dBm output power

Memory

512 kB flash/64 kB SRAM

Transceiver

-96 dBm sensitivity in Bluetooth low energy mode

Pinout

9 GPIO, 8 x 12-bit ADC pins, up to 12 PWM outputs (3 PWM modules with 4 outputs each)
Dimension

2.0" x 0.9" x 0.28" (51mm x 23mm x 8mm)

Supporting OS

Android, iOS, Mac OS X, Windows 8+




3.2 Dispense Mechanism
Table 2 contains design specification for dispense mechanism.

	Table 2. Dispense Mechanism Specification

Item

Specification

Storage dimension of output pill

0.5 cm(l) x 0.5 cm(w) x 1 cm(h)

Digital input

0 or 1 (close or open)

Pill dispensing wait time

5 ± 0.5s




3.3 Phone Case

Table 3 contains design specification for pill dispenser phone case.

	Table 3. Pill Dispenser Phone Case Specification

Item

Specification

Input pill maximum

20

Dimension

Enclosure: 5.836’’ (l) x 3.136’’ (w) x 1.6536’’ (h)

Upper compartment height: 0.8156’’

Lower compartment height: 0.602’’

Material

Plexiglass-Acrylic sheet (thickness=3mm; fluorescent yellow)

Dispense speed

 5 s/pill

Temperature

0 ℃ to 50 ℃



3.4 Samsung Galaxy S6

Table 4 contains product specification for a target smart phone example.
	Table 4. Samsung Galaxy S6 Specification [5]

Item

Specification

Weight

150g / 5.291 oz
Dimension
146.81 mm x 73.41 mm x 8.64 mm / 5.78’’ x 2.89’’ x 0.34’’
OS
Android 5.0.2 Lollipop with TouchWiz
Sensors

Accelerometer, ambient light, barometer, compass, gyroscope, hall, heart rate monitor, HRM, proximity, fingerprint sensor
Battery
3500 mAh



3.5 Android App Specification

Table 5 contains design specification for smart pill phone case android app interface.

	Table 5. Android App Specification

Item

Specification

Function page

3

Connection type

Bluetooth Low Energy

Dimension

Device dependent fullscreen

Additional function

Speaker reminder

Camera flash light reminder

Phone vibration reminder

Fingerprint scanner




4. Design Approach and Details
The Smart Medication Adherence Phone Case consists of two major components: hardware including a pill-dispensing mechanism, an embedded microcontroller and a physical case enclosure; software, which is an Android app, provides a User Interface to implement control on the pill dispenser.
4.1 Hardware
4.1.1 Dispensing Mechanism
The dispensing mechanism of the design consists of two channels that are used to store and dispense pills and two rotation drums driven by motors to transport pills.

[image: image2.png]



Figure 2. CAD design graph showing the dispensing mechanism

The tray can contain 20 normal pills. A rubber band with a pushing stick is used to apply a force pushing the pills towards the rotation drum. 

The rotation drum is special-designed with irregular shape so that the pills coming into it will not collide with the other parts of the structure. And the drum will push the following pills back to the tray so that only one pill can enter the drum at a time. After the motor drives the drum and rotate 180 degrees, the pill will be dropped into another tray and then can be fetched out.
4.1.2 Microcontroller

The controlling circuitry includes an Adafruit NRF52 BLE microcontroller board, two DC motors with corresponding driving circuits, two hall effect sensors and a 2500mAh 3.3V LiPo battery pack. The DC motors are geared down to fit the rotation speed needed for the drums.

[image: image3.png]|

| | N
Adafruit NRF52 Feather Simple amplify stage Motors
I

N M m
o Ul
S5 &
>
Sun®?
w®H
Nl ¥
< =
.
.
.
P R
T eeeee  eeees |1
osf 22233 Wssas o,





Figure 3. System impletation of the controlling circuitry
The driving circuit of the motors are simple BJT-based current amplifiers that amplify the input current of the DC motors.

The microcontroller is programmed in such logic. The BLE board will continuously advertise the BLE signal as peripheral device. Once connected with other Bluetooth controller, it will start recognizing the UART communication sent by the connected device. If some certain characters are received, the motor rotating routine will start running. The hall effect sensors are a simple feedback system to ensure the motor rotates the same degrees every time. A small magnetic is fixed on the drum and the hall effect sensor is placed outside on the cylinder. As soon as the drum rotates to the sensor’s position, the microcontroller will stop the motor from rotating. Thus, although there are some minor delays, the motor can be set to rotate for same degrees every time it rotates.
4.1.3 Enclosure Case

The case enclosure was purposefully designed in 2D in order to quickly prototype. The sketch was drawn and calibrated in AutoDesk’s AutoCAD, and the material was then processed via the laser-cutting machine located at the Invention Studio. In this way, it would be very convenient to update our design when change in dimensions, components or design orientations was necessary.  A grab latch and a hinge were utilized to form the accessible opening area of the case when patients need to manually reload medications.
4.2 Software
All user’s interactions with Medical Adherence Phone Case will be through the mobile App. The functionality specifically designed for the device is integrated into the previous prototype application, BP-n-ME. After the user or medical provider input the treatment and pill-taking schedule, the App is responsible to send push notifications to notify the user to take medicine at exact times. The app will then transmit the signal through BLE to microcontroller with the users’ fingerprint permission to initiate a pill-dispensing action. After receiving the feedback signal from the microcontroller, the App will either record this specific action if it succeeded or report an error if the dispenser fails. A full copy of users’ pill-taking behavior will be stored locally in the app. The updated app menu consisting of “Hardware Status” and “Schedule” subpages is shown in figure 2.

[image: image4.png]Gd BN T 019%E23:18

< Blood Pressure Records

Home
™ Blood Pressure Log

% Medication Utilization
Log

@  Surveys

&« Study Contact

® Call for Clinical Care
Notifications
Hardware Status

®  Schedule




Figure 4. App side menu bar

The hardware status page includes everything users need to interact with the hardware. Functions include making new connections to the BLE hardware, disconnect from the hardware, refill pills and dispense pills after fingerprint authorization. The bottom section of the page monitors and modify the hardware connection status. A green text indicates the connection between the app and phone case is established and users can either make new connection or disconnect from the BLE device. In the top section, users are able to choose from drop down menu which medication to reload or dispense while monitoring the number of pills left in the case.  When reloading, the app requires users to manually input the channel number they just reloaded and keep track of the map of medication to channel. Later on, when the dispense button is pressed, the app will dispense the desired medication from correct channel by sending string bytes to the microcontroller after it asks the fingerprint authentication from users.

[image: image5.png]BN T .418%ME23:20

Hardware Status

Please Select Your Medicine

Acetaminophen

Number of Pills Left

0

DISPENSE PILL

RELOAD PILL

Disconnected

DISCONNECT

CONNECT TO OTHER DEVICES




[image: image6.png]BN 7T .419%H23:18

Enter amount of pills after
reload

CANCEL DONE





Figure 5. Hardware Status Page and Reload Dialog

In the Schedule Page, the user can view his or her pre-existing schedule of the day, including the pill type, the time and the dose. The list is sorted in the timely fashion from the earliest to the latest. The next upcoming pill name and taking time are listed on the top right corner for user’s convenience. The pills can be deleted at any time by clicking the trash bin symbol. Then a dialog will prompt to double verify the delete command. Once deleted the list is refreshed and data is taken out from the database storage.

[image: image7.png]Schedule

Here is your upcoming schedule: v |4:50
Pill Name Time Amount

v 4:50 0 (1]
b 13:50 1 [i]
a 21:50 1 in]
i 22:02 1 [i]
u 22:02 0 in]
y 22:42 0 [i]
t 23:05 0 (1]




[image: image8.png]Add Medicine

Please enter the medicine type below:

Medicine Type

Please choose the dosing time:
PICK A TIME

How many pills:
PICK AN AMOUNT





Figure 6. Schedule Page and Add Medicine Page

On the top right the “plus” symbol will navigate the app to the “Add Medicine” page. The user can click the “back” to discard the information entered or the “save” to add the pill. The page has three rows asking for user information: the name of the pill, the time of taking the pill and the number of the pill taken. Once clicked either of the two navigation buttons on the top, the app navigates back to the Schedule Page and information is either stored or disregarded depending on user’s choice. And the information is synchronized both on the Schedule Page and in the database
4.2.1 Codes and Standards
 
The operating system on the prototype Samsung Galaxy S6 is Android Nougat 7.0, and the development software is Android Studio ver 3.0 which supports:

· Java 8 libraries and Java 8 language features,

· Building separate APKs based on language resources,

· Support for Android Test Support Library 1.0 [10].
4.2.2 Constraints, Alternatives, and Tradeoffs
An alternative design approach would be developing in the iOS platform. One advantage of developing in the iOS platform is unlike the varying versions of Android system running on different devices, the consistent OS version and hardware device will provide stability and prevent potential bugs caused by app running on different devices. However, the Apple Development SDK cost developers more time and money granting the app access of certain sensors of the device such as fingerprints scanner, recording system and camera access. In addition, certain constraints applied during our decision making such as the target users for the phone case is elders who prefer Android phones over iPhones. 
5. Schedule, Tasks, and Milestones
The Medical Adherence team will be implement the system in the following four months. The project milestones and the person assignments as well as the approximate difficulties are being listed in Appendix A. Appendix B contains the GANTT chart displaying the tasks and time management.
6. Final Project Demonstration
For our project, the goal was to successfully simulate the full procedure. In the ideal scenario, such procedure includes:

1. A vibration notification from the “BP-n-ME” app, alerting the patient to take his or her pre-assigned portion of medication;

2. Fingerprint identity check;

3. Successful dispensing of the pre-assigned number of medications from the phone case;

4. Logging of the pill-taking event the “Medication Utilization Log” on the app.

The entire demonstration procedure took place at the Capstone Expo. Dr. Maysam Ghovanloo was present at the Expo and team members performed the Demo on site. 

However, due to the partial success of several features of our hardware/software, simplifications were made to the above-mentioned procedure to ensure initial success in the Demo process. In other words, the currently working procedures of pill-dispensing involves a degree of confinement.

· When installed on new Android devices, The “BP-n-ME”app could sometimes incur crashes when trying to access “Medication Utilization Log” and “Hardware Status” features. As a temporary work-around for the Demo, terminating and restarting the app solved the issue;

· Pills tend to be stuck inside the dispensing cylinder when not oriented vertically. As will be mentioned in Conclusion, this happens due to a combination of the imperfect approach to the design of the dispensing mechanism and inappropriate selection of parts (DC motor is too weak to be powered to spin when a stronger force from rubber band is applied). To address this problem at the Demo, we decided to always orient the case vertically in demonstration (thereof using gravity to dispense pills), disable the rubber band and slow down the DC motor to make spinning more accurate.      

The photos of the completed product are listed in the appendix.

The video link to our group’s Capstone Expo Final Demonstration can be found here on our website: https://www.youtube.com/watch?v=iQPyAJJPPk4
7. Marketing and Cost Analysis
7.1 Marketing Analysis
Currently several products exist in the market, sharing the same goal of increasing medical adherence as the proposed solution. However, the features of the existing products and the technology used behind them are different from the solution. Vitality has made a GlowCap that will glow to remind the patient when it is time for patient to take medication [11], but it doesn’t specify the amount of medication to take and therefore the patients may make mistakes. Another product is the automated medication dispensing service provided by Philips [12], which dispenses the required amount of medication and remind the patient to take them. However, the dispenser is large and therefore is not convenient to carry when the patient is on vacation. The smart medication adherence phone case utilized the sensors build on the smartphone and therefore is smaller in size and cheaper in cost. Thus, the smart medication adherence phone case is unique in the market.
7.2 Cost Analysis
The smart medical adherence phone case is composed mainly of a phone case, the dispensing mechanism inside the case, an Adafruit Feather nRF52 Bluefruit LE microcontroller and an Android App that is used to control. The smartphone doesn’t come with the case and therefore is not counted as a cost. The following table shows a breakdown of the cost of the parts used to make the prototype.

	Item name
	Cost 
	Number


	Total Cost

	
	
	
	

	nRF52
	$25
	1
	$25

	Pill dispenser case
	$30
	1
	$30

	Power Supply, cable and packaging
	$30
	1
	$30

	Total Cost
	$85


Table 6. Component costs for prototype.

The development cost is shown in the following table. The labor cost is assumed to be 30 dollars per hour. The development of the Android App and the experimentation of pill dispensing mechanism is expected to take most of time since their complexity.

	Project Component
	Labour Hour
	Labour Cost
	Part Cost
	Total Component Cost

	
	
	
	
	

	Dispensing Mechanism Design
	10
	$300
	$60
	$360

	Dispensing Mechanism Testing
	40
	$1200
	0
	$1200

	Android App Development 
	50
	$1500
	0
	$1500

	Integration and Testing of All Parts
	20
	$600
	$20
	$620

	Reports
	10
	$300
	0
	$300

	Weekly Meeting 
	200
	$6000
	0
	$6000

	Demo Prep
	10
	$300
	0
	$300

	Total Cost 
	$10280


Table 7. Development Costs.

The fringe benefit is 30% and overhead is 120%. Therefore, the total development cost is $29348, as shown in the table 8.

	Parts
	$85

	
	

	Labor 
	$10200

	Fringe Benefits, % of Labor
	$3060

	Subtotal
	$13340

	Overhead, % of Matl, Labor & Fringe
	$16008

	Total
	$29348


Table 8. Total Development Costs.

The production will consist of 5000 units sold over a 5-year period at a price of $100. Table 4 shows the expected profit and selling price calculation of each product. The sales expense will be 5% of the selling price, which is $12.5. At the price of $250 per unit, the total expected revenue of 5000 units is $1250000, which yields a profit of $38 per unit. The labor cost of assembly and testing is $20 per hour.

	Based on
	5000 Units

	
	

	Parts Cost


	$85

	Assembly Labor 
	$4

	Testing Labor
	$2

	Fringe Benefits, % of Labor
	$1.8

	Subtotal
	$87.8

	Overhead, % of Matl, Labor & Fringe
	$105.36

	Total
	$193.16

	Sales Expenses
	$12.5

	Amortized Development Costs
	$6

	Subtotal, All Costs
	$211.66

	Profits
	$38

	Selling Price
	$250


Table 9. Selling Price per Unit Determination.
8. Conclusion
Currently the pill tray is 60% functional. The smartphone app behaves as expected most of the times, except for some minor bugs when adding pills. However, the hardware components have several issues. First, the approach to the dispensing mechanism is not robust enough. Using pill channels to store pills and depends heavily on the accuracy of the 3D-Printing. A small scratch or bump in the channel or the rotating drum will cause unexpected collision among pills. Secondly, we planned to have a force pushing the pills into the rotating drum. But since our motors have low power consumption, they often fail to rotate when the force is high. The rubber band solution cannot provide a manageable force to push the pills without disabling the motors. Currently we need to make use of the gravity by placing the product upwards. In this case a pill-dispensing success rate of 80% can be reached. For future improvements, the dispensing mechanism might need some redesigning to have a more robust performance. The system voltage can be raised to 5V or higher so that the motors can be replaced by stronger ones with higher rating. Thirdly, the hall effect sensors burned out on the demo, so the motor can only be controlled by rotation time and voltage, which is not accurate enough for continuous pill dispensing. 

The size of the phone case (measured as 5.836’’x 3.136’’x 1.6536’’), which is not small enough to fit into a regular-sized pocket, also provides space for future minor adjustment. One may replace the currently used gear combination in the motor with a smaller one to reduce the size of the design. Moreover, the capacity of the battery being used is beyond the power requirement of the design. Therefore, a smaller battery can be used to further reduce space.

Looking back to our product, the mechanical design took us a large portion of time, but the result is still not as expected. If we could spend more time research on the dispensing mechanism, we should have more preparation over the potential obstacles we are facing. Besides, starting the building-testing-modifying cycle early is another important point when doing such team project. If we have more time, we will improve our dispensing mechanism and the motor control. But since we did not start actual testing early, we had to rush through the testing and improving steps and eventually impact the quality of our product.

During the design expo, the product received a considerably amount of acclaims over the idea of carriable pill tray and the combination of the smartphone and the pill tray case. To improve the overall quality and make it more applicable to market, as stated before, the overall size need to be further reduced by choosing smaller motors and batteries. The dispensing mechanism also need structural enhancement to satisfy the robustness requirement. Moreover, the motors can be upgraded with more complex feedback systems such as rotary encoders or even use the servo motors instead of the DC motors.
9. Leadership Roles
All the tasks assigned to each member are attached in the appendix.

For the team leadership roles, Xiannan Di is the Group Leader that leads all meetings and design progress throughout the semester. Kedeng Pan is the Webmaster that drafts and updates the website for senior design project. Xingyuan Zhu is the Documentation and Parts Coordinator that records project related profiles and organizes the project parts. Yuhao Lin is the Expo Coordinator and is therefore in charge of expo demonstration details. Zhipeng Shao is the Email Manager that drafts weekly email report.
10.  References
[1] IMS Institute for Healthcare Informatics Avoidable costs in US health care. 2013. Available at:http://www.imshealth.com/deployedfiles/imshealth/Global/Content/Corporate/IMS%20Institute/RUOM-2013/IHII_Responsible_Use_Medicines_2013.pdf. [Accessed 30 Oct. 2017].
[2] Ho, P., Bryson, C. and Rumsfeld, J. (2009). Medication Adherence: Its Importance in Cardiovascular Outcomes. Circulation, 119(23), pp.3028-3035.

[3] Comstock, J. (2017). Slideshow: 8 pillboxes that connect to your phone. [online] MobiHealthNews. Available at: http://www.mobihealthnews.com/20795/slideshow-8-pillboxes-that- connect-to- your-phone/page/0/4 [Accessed 30 Nov. 2017].

[4] Philips Lifeline, "Automated Medication Dispensing Service," Philips, 2012. [Online]. Available: https://www.lifeline.philips.com/pill-dispenser/health-mdp.html. [Accessed 1 December 2017].

[5] MARTONIK, A. (2017). Samsung Galaxy S6 active specs. [online] Android Central. Available at: https://www.androidcentral.com/samsung-galaxy-s6-active-specs [Accessed 30 Nov. 2017].

[6] Nordic. (2017). nRF52 DK Specification. [online] Available at: https://www.adafruit.com/product/3406 [Accessed 24 Jan. 2018].

[7] Nordic. (2017). Bluetooth 5 ICs/solutions. [online] Available at: https://www.nordicsemi.com/eng/Products/Bluetooth-5 [Accessed 2 Dec. 2017].

[8] Universal Serial Bus Cables and Connectors Class Document. (2017). 7th ed. [ebook] USB Cables and Connectors Class Document, pp.30-69. Available at: http://www.usb.org/developers/docs/devclass_docs/CabConn20.pdf [Accessed 2 Dec. 2017].


[9] Arm Mbed. (2017). Mbed OS | Mbed. [online] Available at: https://www.mbed.com/en/platform/mbed-os/#features [Accessed 2 Dec. 2017].

[10] Notes, A. (2017). Android Studio Release Notes | Android Studio. [online] Developer.android.com. Available at: https://developer.android.com/studio/releases/index.html [Accessed 2 Dec. 2017].

[11]"Vitality - NantHealth", NantHealth, 2017. [Online]. Available: https://nanthealth.com/vitality/. [Accessed: 25- Oct- 2017].

[12] "Medication Dispensing Service | Philips Lifeline ®", Philips, 2017. [Online]. Available: https://www.lifeline.philips.com/pill-dispenser/health-mdp.html. [Accessed: 30- Nov- 2017]

.
Appendix A
Project Milestones Table
Risk levels are evaluated by relevant previous background or experience. High risk level means no previous experience is presented.
	Task Content
	Member Assignment
	Risk Level

	Planning, Preparation and Deliverables
	
	

	  Order Parts
	All
	Low

	  Project Proposal and Summary
	All
	Medium

	  Proposal Presentation
	All
	Low

	  Progress Report
	All 
	Low

	  Final Demonstration and Presentation
	All
	Medium

	  Final Report
	All
	High

	Pill Tray Case Design
	
	

	  Dispensing Mechanism and 3D Printing
	All
	High

	  Microcontroller Control and Power Distribution
	Zhu, Lin
	Medium

	  Case Design 
	Lin, Zhu
	Medium

	Smartphone App Development
	
	

	 Bluetooth Communication and Hardware Status
	Pan
	High

	  UI Design
	Pan, Di
	Medium

	  App Function
	Di
	High

	App and Device Communication
	
	

	  Bluetooth Communication
	Lin, Pan, Di
	Medium

	Inspection and Improvement
	All
	Low


Appendix B
GHANTT Chart
[image: image9.png]F o = e — = = = = =
o P= T Lo | Lo L
B s o T e e e
E IS R - J—
[ 2 v won s wes
T b e on Wt e wa
o b
C s Samamem wem e neves o
©m oi— o o T
Tm e mem st e
[0 4 mrcume e e s
pmy
o e wen i e
R T -
I
= e o Sr A e
-
[ =
W e mem e
Bm e mem e e
W i wen s s
T4 e e s e
e wem e
—
F i =
P e Tam e e
= ven wesmsn
P vt en W Teawm
fnty o
PR U R
b s can e o
b o e e ow
P - o
pe— = — —— ———
= ¢ == Pt g o -
ot





Appendix C
Product Photos
Can be found in the provided link:

http://ece4012y2018.ece.gatech.edu/spring/sd18sGM1/
[image: image10.jpg]


[image: image11.jpg]


[image: image12.jpg]


[image: image13.jpg]


[image: image14.jpg]


[image: image15.jpg]


[image: image16.jpg]



Appendix D

App and microcontroller codes, CAD files can be found on the website

http://ece4012y2018.ece.gatech.edu/spring/sd18sGM1/
2

